Supplementary Tables and Figures

Room temperature roll-to-roll additive manufacturing of polydimethylsiloxane-based centrifugal microfluidic device for on-site isolation of ribonucleic acid from whole blood

# 1,3Trung Hoang, 1,3,5Han Truong, 1,3,5Jiyeon Han, 2Saebom Lee, 2Jihyeong Lee, 4,5Sajjan Parajuli, 1,2Jinkee Lee*, 1,3,4,5GyouJin Cho*

# 1Institute of Quantum Biophysics, Sungkyunkwan University, Suwon, Korea.

# 2School of Mechanical Engineering, Sungkyunkwan University, Suwon, Korea

# 3Department of Biophysics, Sungkyunkwan University, Suwon, Korea

# 4Research Engineering Center for R2R Printed Flexible Computer, Sungkyunkwan University, Suwon, Korea

# 5Department of Intelligent Precision Healthcare Convergence, Sungkyunkwan University, Suwon, Korea

# *Corresponding author:

Name: Jinkee Lee, GyouJin Cho

Address: Institute of Quantum Biophysics, Department of Biophysics, Sungkyunkwan University, Suwon, Republic of Korea
E-mail: lee.jinkee@skku.edu; gcho1004@g.skku.edu

a

b

Fig S1. (a) A detailed structure of the roll-to-roll imprinting unit. 
               (b) Schematic images of step-by-step manufacturing process of PDMS based LoaD devices.

Fig S2. Products manufactured by R2R systems. (a) PET substrates before & after treated with adhesion promoter by gravure coating methods; (b) Imprinted chip on PET substrates.

Fig S3. Design of plasma separation chamber on centrifugal microfluidic devices. (a) 3D design of plasma separation chamber with different geometries; (b) Plasma separation efficiency measurement by experimenting on different chamber designs.

Fig S4. Demonstration of the actuation mechanism of ferro-wax valves integrating on our chip devices.

Table S1. Autofluorescence intensity of glass and three PDMS formulations at four different excitation wavelengths.

|  | DAPI | FITC | TRITC | CY5 |
| --- | --- | --- | --- | --- |
| Glass | 388.16 ± 0.1 | 398.64 ± 0.15 | 391 ± 0.11 | 398.17 ± 0.06 |
| Sylgard 184 | 389.57 ± 0.01 | 401.94 ± 0.1 | 392.02 ± 0.04 | 396.97 ± 0.05 |
| Vinyl-terminated PDMS | 390.56 ± 0.04 | 392.88 ± 0.18 | 390.44 ± 0.32 | 396.7 ± 0.06 |
| Room-temp PDMS | 389.78 ± 0.16 | 399.91 ± 0.13 | 391.76 ± 0.12 | 393.13 ± 0.07 |

Fig S5. Water contact angle measurements of PDMS samples.

Table S2. Optimization of R2R imprinting parameters

| Applied nip Pressure (kgf/cm2) | Web transfer speed (mm/s) | Roll temperature (oC) | Humidity (%) | Replication accuracy |  |
| --- | --- | --- | --- | --- | --- |
|  |  |  |  | Height | Width |
| 1 | 0.3 | 25 | 45 | 99% | 99% |
| 4 | 0.3 | 25 | 45 | 71.7% | 92.50% |
| 10 | 0.3 | 25 | 45 | 69.5% | 92% |

Table S3. Operation steps for RNA isolation on R2R imprinted lab-on-a-disc

|  | Step | Rotational speed (rpm) | Time(s) | Operation |
| --- | --- | --- | --- | --- |
| 1 | Plasma separation | 2000 | 180 |  |
| 2 | Lyse | - | 10 | Valve 1 open |
| 3 |  | 1000 | 30 | Channel mixing (Plasma, Binding buffer, Proteinase K) |
| 4 |  | - | 10 | Valve 2 close |
| 5 | Binding | - | 10 | Valve 3 open (Ethanol) |
| 6 |  | 1000 | 15 | Transfer ethanol to mixing chamber |
| 7 |  | Mix | 30 | Chamber mixing |
| 8 |  | - | 600 | Incubate at 600C |
| 9 |  | - | 10 | Valve 4 open |
| 10 |  | 1000 | 15 | Transfer magnetic nanobeads to mixing chamber |
| 11 |  | Mix | 60 | Mixing lysate mixture with magnetic nanobeads and incubate at room temperature (RT) |
| 12 |  | - | 10 | Valve 5 open (waste chamber) |
| 13 |  | 1000 | 15 | Transfer the aqueous mixture to waste chamber |
| 14 | Wash | - | 20 | Valve 5 close (waste chamber), Valve 6 open (VWM1) |
| 15 |  | 1000 | 15 | Transfer VWM1 buffer to mixing chamber |
| 16 |  | Mix | 30 | Chamber mixing |
| 17 |  | - | 10 | Valve 5 open (waste chamber) |
| 18 |  | 1000 | 15 | Transfer VWM1 buffer to waste chamber |
| 19 |  | - | 180 | Repeat 13-17 process with RWA2, WE buffer. |
| 20 | Elute | - | 10 | Valve 5 close (waste chamber), Valve 9 open |
| 21 |  | 1000 | 15 | Transfer ER buffer to mixing chamber |
| 22 |  | Mix | 30 | Chamber mixing |
| 23 |  | - | 10 | Valve 10 open (Elution chamber) |
| 24 |  | 1000 | 15 | Transfer solution to elution chamber |

Fig S6. Failure of tensile measurement on vinyl terminated PDMS samples due to low modulus. (a) Comparison of three types of PDMS after gripping on tensile measurement machine; (b) Crack formation of vinyl-terminated PDMS during gripping process.

Fig. S7. Snapshot images of the top surface design of centrifugal microfluidic devices.

Fig. S8. Reactor systems for storing PDMS mixture during the R2R imprinting operation

Fig. S9. The centrifugal machine made by Optolane (Korea) company.

Fig. S10. Design and optimization of current flexible mold space for enhancing production rate.